## Slide 1: High-throughput MR-IGRT of orthotopic pancreatic tumours in mice using the SARRP
Danny Allen, Stuart Gilchrist, Veerle Kersemans, Paul Kinchesh, Amy Elliott, James Thompson, Sophie Hughes, Simone Lanfredini, Asmita Thapa, Eric O'Neill, Mark Hill, Sean Smart
Preclinical Imaging SRF - Radiation Biophysics SRF - Cell Signalling Group
Oxford Institute for Radiation Oncology

## Slide 2: Why MR-IGRT ?
CT
CT in the mouse offers poor contrast of most soft tissues, including pancreatic tumours.
Tumour identification and RT planning not always possible.

## Slide 3: Why MR-IGRT ?
MR
MRI offers good contrast of many:
  organs
  tumours
  tissue functions

## Slide 4: Why MR-IGRT ?
MR
CT

## Slide 5: Why MR-IGRT ?
Tumour
MR
CT

## Slide 6: Prior art: MR-IGRT in Oxford
In collaboration with our Mechanical Workshop and the Advanced Technology Group (Vojnovic), we have previously validated the use of MR-IGRT for targeting of:
  Adrenal glands [1]
  Genetically-invoked pancreatic tumours [1]
  Inward-growing xenografts [2]
[1] Kersemans V, et al. (2017) An efficient and robust MRI-guided radiotherapy planning approach for targeting abdominal organs and tumours in the mouse. PLoS ONE 12(4): e0176693. https://doi.org/10.1371/journal.pone.0176693
[2] Corroyer-Dulmont A, et al. (2017) MRI-guided radiotherapy of the SK-N-SH neuroblastoma xenograft model using a small animal radiation research platform. The British Journal of Radiology 90:1069

## Slide 7: Prior art: MR-IGRT in Oxford
But this developmental workflow required about 6 postdocs; not economically viable
The workflow has been refined so it works at high-throughput in the hands of non-expert users
Here we describe the optimisations starting with the lab floorplan

## Slide 8
MRI
SARRP
MRI
MRI
CT/PET/SPECT
Ivis
US
Microscope

## Slide 9
MRI
SARRP

## Slide 10
MRI
SARRP
MR scanner and SARRP are in different rooms but within the same animal facility

## Slide 11
MRI
SARRP
MR and CT compatible cradles with mouse must be transferred between MRI scanner and SARRP

## Slide 12: Technical Requirements
MR and CT compatible animal handling rigs
Fast, motion artefact free MRI scans
Simple tumour identification
Simple tumour segmentation

## Slide 13: Technical Requirements
Robust image data transfer mechanism
Robust MR->CT registration software
Easy reporting and upload of spatial parameters for SARRP RT plan
  Including collimator selection

## Slide 14: MR-SARRP Cradle

## Slide 15: MR-SARRP Cradle
Cradle must allow:
  Good positioning and immobilisation
  Anaesthesia and heating
  High-resolution, distortion-free imaging
  Quick, reliable transfer between systems

## Slide 16: MR-SARRP Cradle
Use low atomic number materials (CT ✔︎)
Electrical heating via high frequency AC (MR ✔︎) through a carbon fibre (CT ✔︎) resistor [3]
Customised adaptors for cradle positioning in:
  Bruker and Varian MRI scanners
  SARRP
[3] Kersemans V. et al. A Carbon-Fiber Sheet Resistor for MR-, CT-, SPECT-, and PET-Compatible Temperature Maintenance in Small Animals. Tomography. 2019 Jun;5(2):274-281.

## Slide 17: MR-SARRP Cradle
3D print main components (10 pounds).
carbon fibre sheet resistor (50 pence)
plastic tubing with Luer fitting for gas supply (50 pence)
pneumatic monitor of breathing (5 pounds)
non-magnetic D-connector for electrical services
coupled to existing physiological monitor system on MRI and SARRP (10K each)

## Slide 18: MR-SARRP Cradle
Tooth bar
Gas tube
Cradle
Carbon fibre heater
Cradle Adapter
Cable
Gripper

## Slide 19: MR-SARRP Cradle

## Slide 20
Cradle
Anaesthetic
Pressure Balloon
Fibre-optic
thermometer
Heating

## Slide 21
Cradle positioned in the MRI scanner

## Slide 22
Cradle being transferred

## Slide 23
Cradle positioned in the SARRP

## Slide 24: MR Scan Modes

## Slide 25: MR Scan Modes
Respiratory gated scans that operate at constant TR:
    Arbitrary choice of TR
      Minimum scan times
      Maximally efficient scanning
Use separate image sets for:
  Tumour identification (high contrast, low res, fast)
  MR-CT registration procedures (low contrast, high res, slow)

## Slide 26: Tumour identification
[4] Kinchesh P. et al. Reduced respiratory motion artefact in constant TR multi-slice MRI of the mouse. Magnetic Resonance Imaging 60 (2019) 1–6
Combined T2 and T1-weighted FSE [4]
Anisotropic resolution (0.2x0.2x0.5 mm)
Not ideal for image registration
Fast: <90s per volume
Good tumour contrast

## Slide 27: Tumour segmentation (ITK-SNAP)

## Slide 28: MR-CT registration
Resp-gated bSSFP [5]
Isotropic resolution (0.2x0.2x0.2mm)
Good for registration
Takes ~6 minutes
Some contrast of the tumour is lost.
[5] Gomes A. et al. Cardio-Respiratory synchronized bSSFP MRI for high throughput in vivo lung tumour quantification. PLoS ONE 14(2): e0212172. https://doi.org/ 10.1371/journal.pone.0212172

## Slide 29: MR / CT registration
mr - fse
mr - bssfp
CT
Tumour segmentation based on mr fse image

## Slide 30: MR / CT registration
mr - fse
mr - bssfp
CT
Find transformation from bssfp to CT using registration

## Slide 31: MR / CT registration
mr - fse
mr - bssfp
CT
Transformation from fse to bssfp is known (same scanner)

## Slide 32: MR / CT registration
mr - fse
mr - bssfp
CT
fse and segmentation can now be transferred to CT coordinates

## Slide 33: MR / CT registration
mr - fse
mr - bssfp
CT

## Slide 34: Data Transfer

## Slide 35
MRI
SARRP

## Slide 36
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
3 computers are involved in this process:

## Slide 37
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
MRI Acquisition

## Slide 38
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
Tumour Segmentation

## Slide 39
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
MRI -> CT Registration

## Slide 40
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
Robust and efficient data transfer needed between these computers

## Slide 41
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
A shared folder on the MRI PC is mounted on the SARRP and Biopac PC to facilitate file exchange

## Slide 42
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
The first few MRI scans are positioning

## Slide 43
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
The first few MRI scans are positioning

## Slide 44
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
The first few MRI scans are positioning

## Slide 45
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
When the fse scan is acquired it needs to be copied to the shared folder

## Slide 46
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
To avoid operator error this is done using a script called by a GUI

## Slide 47
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v

## Slide 48
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
v
As soon as the fse file is visible on the Biopac PC the tumour can be segmented and the segmentation mask saved to the shared folder

## Slide 49
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
v
Meanwhile the bssfp scan has been acquired ...

## Slide 50
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
... and is copied to the shared folder using the transfer script/GUI

## Slide 51
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg

## Slide 52
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
MRI scanning is now over and the mouse is transferred to the SARRP where the CT image is acquired

## Slide 53
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
bssfp
fse
shared folder
seg
v
fse
CT
SARRP study
bssfp
seg
All of the relevant scans are now visible from the SARRP PC. The registration software is started and automatically moves the fse,bssfp and seg files from the shared folder to the local SARRP study folder

## Slide 54
scan 3
scan 4
scan 2
scan1
scan 5
SARRP PC
BIOPAC PC
MRI PC
MRI study
v
shared folder
v
shared folder
v
shared folder
v
fse
CT
SARRP study
bssfp
seg
The shared folder is now empty and this tells the MRI operator that they can transfer the next set of data

## Slide 55: MR/CT Registration

## Slide 56
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 57
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 58
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 59
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen
mr-bssfp
CT

## Slide 60
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 61
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 62
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration [1]
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen
[1] Heinrich MP, Jenkinson M, Bhushan M, Matin T, Gleeson FV, Brady SM, et al. MIND: modality independent neighbourhood descriptor for multi-modal deformable registration. Med Image Anal. 2012; 16 (7):1423–35.

## Slide 63
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 64
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 65
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 66
User crop CT
Rigid registration
User check Y-axis registration
User crop MR
Deformable registration
fse+seg -> bssfp -> CT transformation
Bounding-sphere/cross-hairs superimposed on fse image
Transformed annotated fse write to DICOM
Tumour bounding sphere and centre coordinates output to screen

## Slide 67
Annotated registered MR loaded into SARRP RT planning software
Tumour position determined from registered MR alone
Dose planning based on CT

## Slide 68: Personnel

## Slide 69
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC

## Slide 70
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC

## Slide 71
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
MR Operator:
Anaesthetize animal
Place animal in scanner
Run MRI scans
Transfer MRI data to SARRP

## Slide 72
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
SARRP Operator:
Run SARRP CT Scan
Run MR/CT registration software
SARRP Dose planning
Deliver RT

## Slide 73
MRI
SARRP
MR PC
SARRP PC
BIOPAC PC
Director:
Transfer animal
Place animal in SARRP
Monitor animal (BIOPAC)
Segment Tumour
Direct dose planning

## Slide 74
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
Time
SARRP Operator
Director
Time
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
MR Operator
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 75
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
Time
Time
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 76
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 77
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 78
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 79
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 80
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 81
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 82
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse

## Slide 83
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
Transfer Mouse
Transfer Mouse

## Slide 84
Time
Time
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
CT Scan
MR->CT Registration
Write MR to DICOM
Plan RT
Deliver RT
MR Operator
Anaesthetize mouse
Position in Scanner
MRI-fse
MRI-bssfp
Segmentation
Segmentation
Transfer Mouse
Plan RT
Remove Mouse
Transfer Mouse
Transfer Mouse
20-30 minutes
Turnaround time

## Slide 85: The Study

## Slide 86
Orthotopic pancreatic tumour (500 NKPCB NBC12.1f cells in 5 µl matrigel)

## Slide 87
Orthotopic pancreatic tumour (500 NKPCB NBC12.1f cells in 5 µl matrigel)
8 Mice per group

## Slide 88
Orthotopic pancreatic tumour (500 NKPCB NBC12.1f cells in 5 µl matrigel)
8 Mice per group
8 groups

## Slide 89
4 groups MR-IGRT

## Slide 95
        ~100mm3
Fractionated RT (4Gy)
On days 1-3-5 (mon-wed-fri)
Follow up MRI at 1,3,5 weeks using fse at 4.7T
Schedule 1
Tissue harvest
Post mortem analysis

## Slide 96
        ~100mm3
Fractionated RT (4Gy)
On days 1-3-5 (mon-wed-fri)

## Slide 97
Mon | Tue | Wed | Thu | Fri | Sat | Sun
29 | 30 | 1 | 2 | 3 | 4 | 5
6 | 7 | 8 | 9 | 10 | 11 | 12
13 | 14 | 15 | 16 | 17 | 18 | 19
20 | 21 | 22 | 23 | 24 | 25 | 26
27 | 28 | 29 | 30 | 31 | 1 | 2
3 | 4 | 5 | 6 | 7 | 8 | 9
3 Imaging sessions per week for 4 weeks
Morning session: 9am-1pm
8 Mice per session
96 MR-IGRT planned in total
All delivered on schedule
May
June
April/May/June 2019

## Slide 98: Benefits
Scientific optimisation and clinical relevance
Low animal welfare cost
Production line workflow gives economic viability

## Slide 99: Conclusion
High throughput MR-IGRT is available as a routine tool for treatment of abdominal tumours
imagingcore@oncology.ox.ac.uk